Table 1. Determination of PARK7/DJ-1 protein in human blood serum and synthetic cerebrospinal fluid samples (n=3).

| Sample | Added (µg L−1) | Found (µg L−1) | Recovery (%) |
| --- | --- | --- | --- |
| Blood serum | 30.0 | 31.2±3.6 | 104 |
|  | 40.0 | 37.6±3.4 | 94.0 |
|  | 100 | 103±5 | 103 |
| Cerebrospinal fluid | 30.0 | 30.4±2.4 | 101 |
|  | 40.0 | 41.1±1.2 | 103 |
|  | 100 | 101±3 | 101 |

Table 2. Analytical performance of different immunosensor platforms for the diagnosis of Parkinson’s disease.

| Sample | Immunosensor | Immobilization | Target | LDR | LOD | Stability | Ref. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Blood plasma | GCE voltammetric immunosensor | Anti-α-Syn (EDC/NHS couple) | α-synuclein | 0.05−500 fg mL−1 | 0.03 fg mL−1 | 7 days | [33] |
| Blood serum | FTO voltammetric and impedimetric immunosensor | Anti-α-Syn (cystamine SAM) | α-synuclein | 10−1000 ng mL−1 | 1.13 and 3.62 ng mL−1 | 7 days | [34] |
| Blood serum | SPE impedimetric immunosensor | Anti-DJ-1 (cysteamine- glutaraldehyde binding) | DJ-1 protein | 40−150 ng mL−1 | 7.5 ng mL−1 | − | [31] |
| Cerebrospinal fluid and saliva | ITO voltammetric and impedimetric immunosensor | Anti-DJ-1 (SAM) | DJ-1 protein | 4.7–4700 fg mL−1 | 0.50 fg mL−1 | 5 weeks | [35] |
| N2A cells | MIP/SPCE voltammetric immunosensor | Anti-DJ-1 (SAM) | DJ-1 protein | 1.0−500 nmol L−1 | 1.0 nmol L−1 | − | [36] |
| Blood serum and cerebrospinal fluid | 3D-printed voltammetric and impedimetric immunosensor | Anti-DJ-1 (EDC/NHS couple) | DJ-1 protein | 5.0–500 µg L−1 | 1.01 and 2.69 µg L−1 | 25 days | This work |