Table 1. Electrochemical characteristics for 3D printed sensors in presence of 1.0 mmol L-1 [Fe(CN)6]3−/4− obtained from cyclic voltammetric experiments.

| Conductive filament | Ipa / µA | Ipc / µA | Ipa / Ipc | ΔEp / mV |
| --- | --- | --- | --- | --- |
| G-PLA | 0.54 | -0.17 | 3.18 | 869 |
| G-PLA treated* | 15.10 | -13.50 | 1.12 | 260 |
| CB-PLA | 5.67 | -4.18 | 1.36 | 904 |
| CB-PLA treated* | 12.20 | -11.00 | 1.11 | 466 |
| Gpt-PLA | 19.80 | -18.90 | 1.05 | 264 |
| Gpt-PLA treated* | 20.50 | -20.60 | 0.995 | 267 |

* Electrochemical pre-treatment: application of +1.4 V for 200 s, followed by the application of -1.0 V for 200 s, in presence of 0.5 mol L-1 NaOH.

Table 2. Parameters obtained from the electrochemical characterization with the proposed and commercial filaments.

| Conductive filament | k0obs x10-6 / cm s-1 |  | ΔEp / V | Rct / Ω | Electrochemically active surface area / |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
|  | CV | EIS |  |  | mm2 % |  |  |
| G-PLA treated* | 6.1 | 10.1 | 0.196 | 26.4 | 10.26±0.07 | 81.6±0.5 |  |
| CB-PLA treated* | 5.0 | 9.1 | 0.166 | 29.4 | 10.26±0.06 | 81.6±0.5 |  |
| Untreated Gpt-PLA | 7.5 | 14.6 | 0.173 | 18.2 | 16.72±0.03 | 133.0±0.2 |  |

* Electrochemical pre-treatment: application of +1.4 V for 200 s, followed by the application of -1.0 V for 200 s, in presence of 0.5 mol L-1 NaOH.

Table 3. Analytical features for the detection of UA using Gpt-PLA 3D printed electrode.

| Characteristics | Values |
| --- | --- |
| Linear range / µmol L-1 | 0.5 to 150.0 |
| R2 | 0.998 |
| Sensitivity / µA µmol-1 L | 0.05 |
| LOQ / µmol L-1 | 0.23 |
| LOD / µmol L-1 | 0.07 |
| Intra-electrode (RSD, n = 13) | 4.8 |
| Inter-electrode (RSD, n = 5) | 12.0% |

Table S1. Optimization of DPV parameters for UA detection.

| Parameters | Range | Optimal value |
| --- | --- | --- |
| Modulation time / ms | 5 – 50 | 30 |
| Step potential / mV | 1 – 10 | 80 |
| Modulation amplitude / mV | 10 – 100 | 5 |